## Slide 1: Assuring Quality in POCT
Dr Adilahtul Bushro binti Zaini
Medical Microbiologist
Hospital Sg Buloh
National POCT Committee

## Slide 2: Introduction
POCT has important role to play in diagnosis, treatment and monitoring of infectious disease
In a larger scale POCT has prove itself as a very important role in control and mitigate pandemic in particular.

## Slide 3: Flow of process following POCT testing
As for many POCT in infectious disease, the result in needed as urgent for example Dengue RTK, PCT.
Others maybe semiurgent for example HIV antibody, HCV and Syphilis antigen
It serve as a screening tool before proceed to lab based confirmatory test.
There is a need for complete algorithm including patient selection (preanalysis), analysis and post analysis before embark on POCT.
Example-Hep C & HIV RTK included in national diagnostic algorithm.
As for syphilis antigen, the algorithm need to be informed and training would be very helpful

## Slide 4: POCT for Infectious disease pathogen in Malaysia
COVID antigen RTK
Dengue Combo RTK
HIV antibody RTK
HCV antigen RTK
Syphilis antigen RDT

## Slide 5: Quality Assurance
A program to assured the quality of POCT test that is being provided
All party providing the test is responsible for assuring the quality start from the policy maker and stake holder to the end user.
In central lab where the test is highly sensitive/accurate and in controlled environment, there is also program to assure quality, whatmore in a periphery where the turnover rate of patient is high.

## Slide 6
Start from the Evaluation of test method
Acceptance Test
Method Verification
IQC(internal quality control)
EQC
Audit & Assessment

## Slide 7: Evaluation of the test method
During the market survey, potential candidates can send a sample of kit for evaluation by designated laboratory
Done at designated laboratory e.g. IMR,MKAK
Sample is tested on a known sample(positive& negative) and report on sensitivity/specificity is produced.
Based on specification set by the tender committee, performance evaluation report should be included.
Only kit that passed the requirement will be considered in technical committee or any type of purchasing(direct purchase etc).
Technical committee will give recommendation based on the kit performance and other criteria

## Slide 8: Acceptance test
Every batch of production is different. There may be lot to lot variation. Therefore there must be quality check on this.
Manufacturer usually have their own QC check before release the batch. However, there are still who don’t have that good manufacturing practice. Example: recent DCTK manufacturer.
The best is that manufacturers are ISO accredited.
Another way to ensure a quality product is received is to ask the supplier to send for evaluation at designated laboratory in every shipment or at regular interval. Example: testing at every consignment received (HIV).
This can avoid a faulty kit distributed and much more difficult to recall.

## Slide 9: Method Verification
Although the test kit already approved by the tender committee for purchasing, testing at end user is still mandatory to ensure it meets the intended use.
The kit may be of different lot, different shipment from the evaluation. In addition, the facility condition and local personel may have the impact on the kit performance. By doing method verification, the exact performance of the kit is known. In this process, the test is run with known sample and repeatedly with different user. Report on accuracy and reproducibility is produced.
Method verification can be done either at selected site or all site and the vendor should help in this process.
Method verification done at Pathology Department for DCTK able to detect low sensitivity test kit, markedly different from Evaluation report during tender process (30% vs 95% that claimed by manufacturer)

## Slide 10: Monitoring-IQC
The frequency ICQ mentioned in National POCT Policy & Guideline
Divided to High, Moderate, Low complexity
Immunochromatographic test(ICT)/ Lateral flow Assay(LFA) considered low complexity test.
MONTHLY testing is a minimum requirement. Testing with positive & negative control(1 respectively)
Ideally the box should comes with positive and negative LIQUID CONTROL. However, this will incur higher cost.
Alternatively, testing with known sample is ACCEPTABLE.

## Slide 11
Another option is control material is included in every batch lot only and done at designated testing site. Vendor should help in this process (put in specification). Example: Glucometer IQC has ICQ provided by manufacturer for all users.
Testing for control band is not sufficient for performance check.
This process should be oversight by a laboratory personnel as a gatekeeper for quality in testing.
Example there was a case where testing of tap water on glucometer give reading of 3-4mmol/L

## Slide 12: EQC-External Quality Control
External quality assessment (EQA)program, a group of laboratories analyze the same specimens and submit their results to a central facility where the data are examined for outliers, it is also known as proficiency test
Proficiency test (PT) is mandatory in quality monitoring
Quality of reagent and competency of personnel are being evaluated simultaneously.
PT is provided by ISO 17043 accredited provider. Example RCPA(Australia), API (USA) or KKM(IMR/MKAK)
This requirement should be included in a tender specification.
Example: HIV antibody test(IMR provided samples for 5 KK and Hospital).
DCTK: Testing the kit with RCPA samples detect low sensitivity of DCTK

## Slide 13: Other element for quality for POCT
Document control- Local procedures, processes and record retention guidelines should be followed. Original results and evaluation approval should be kept at one location, and all locations using the device are to have a copy and/or access to the documentation for referencing at any time.
Training-All POCT users must receive training for the POCT device they will be using. Users must be deemed competent to perform the testing before being granted access.

## Slide 14
Audit-The most significant risk associated with POCT is the failure of clinical users to follow POCT policies and procedures. This can result in patient safety issues and jeopardize program compliance with accreditation standards.
Quality indicator monitoring

## Slide 15: Conclusion
POCT has important role to provide rapid and simple test for efficient infectious disease management
However, the test can be as good as its get when it is properly planned and executed.
Quality Assurance Programme should be in every policy maker consideration because the test was taken away from the central lab where the environment is controlled and properly monitored.

## Slide 16: Thank You